# Supplementary Information for

# Microfluidic multiple cross-correlated Coulter counter for improved particle size analysis

Wenchang Zhanga,1, Yuan Hu a,1, Gihoon Choib,

Shengfa Lianga, Ming Liua*, and Weihua Guanb, c*

a Key Lab of Microelectronic Devices & Integrated Technology, Institute of Microelectronics, Chinese Academy of Sciences, Beijing 100029, China

b Department of Electrical Engineering, Pennsylvania State University, University Park 16802, USA

c Department of Biomedical Engineering, Pennsylvania State University, University Park 16802, USA

*Corresponding Authors: liuming@ime.ac.cn (M. Liu), w.guan@psu.edu (W. Guan)

# Derivation of multiple cross-correlation analysis

For microfluidic Coulter counter with n sensing pores, the current dip could be modeled with Gaussian function with a DC baseline IDC:

(s-1)

where ΔIi (i = 1, …, n) is constant representing the heights of the current dip; σ is the pulse width; τi-1 is the time-delay between f1(t) and fi(t). Assuming a constant particle flow velocity, τi-1 = (i-1)τ1. The conventional cross-correlation analysis could be extended as follow:

(s-2)

Traditionally, MDC is related to IDC and pulse width without providing information about the time-relative between multiple pulse signals. Therefore, MDC is removed and only MAC(τ) is used for analysis. Meanwhile, just the maximum value of MAC(τ) can provide useful information, there is no need to calculate the complete cross-correlation. MAC(τ) reaches its maximum value when τ = τ1. Thus, the maximum value (denoted by M) can be calculated directly when the time-delay τ1 is known, as follow:

(s-3)

Therefore, the M with different values of n could be calculated:

n = 2,

(s-4)

n = 3,

(s-5)

n = 4,

(s-6)

n = 5,

(s-7)

Based on the above derivation, the M with different n could be generalized. The arithmetic expression could be simplified with the assumption that the heights of the n pulse peaks are equal (ΔI1 = ΔI2 = …= ΔIn):

(s-8)

Where  is the number of k-combinations.

Meanwhile, for the microfluidic chip designation in this work, the value of IDC is a few hundred nanoamperes, while that of ΔI ranges from a few tenths of a nanoampere to a few nanoamperes. Therefore, the dominant component (k = 1) could be used to approximate the value of M:

(s-9)

As assumed in the manuscript,, the expectation and the variance. According to the statistical analysis theory, we conclude that:

(s-10)

Therefore, .

In order to compress the distribution interval of M to span the same range as ΔI, the probability distribution density function should be used for analysis.

The probability distribution density function of ΔI and M could be expressed as:

(s-11)

According to equation (s-9), fM(M) could be transformed,

(s-12)

Therefore, after compression, the probability density of M distributes at the same middle location as that of ΔI (, ); the standard deviation of M (,) is shrunk down to of that of ΔI.

Figure S1. Single sized bead characterization. (A) Time series of the current profile measured within 10.5 s with 6 μm polystyrene beads passing through. (B) One of the enlarged five pulse signals of the profile circled in (A).  (C) Histogram of ΔI of 9,550 events (6 μm). (D) Histogram of M of 9,550 events (6 μm). (E) Time series of the current profile measured within 10.5 s with 8 μm polystyrene beads passing through. (F) One of the enlarged five pulse signals of the profile circled in (E).  (G) Histogram of ΔI of 9,800 events (8 μm). (H) Histogram of M of 9,800 events (8 μm). (I) Time series of the current profile measured within 10.5 s with 8.5 μm polystyrene beads passing through. (J) One of the enlarged five pulse signals of the profile circled in (I). (K) Histogram of ΔI of 8,946 events (8.5 μm). (L) Histogram of M of 8,946 events (8.5 μm). (M) Time series of the current profile measured within 10.5 s with 9 μm polystyrene beads passing through. (N) One of the enlarged five pulse signals of the profile circled in (M).  (O) Histogram of ΔI of 9,487 events (9 μm). (P) Histogram of M of 9,487 events (9 μm). (Q) Time series of the current profile measured within 10.5 s with 10 μm polystyrene beads passing through. (R) One of the enlarged five pulse signals of the profile circled in (Q).  (S) Histogram of ΔI of 8,632 events (10 μm). (T) Histogram of M of 8,632 events (10 μm).